## Sheet1
Log of all Optimass 6400 Two-phase tests. Includes date, test number, nominal flow rate and temperature set point. Tests last up to half a day. Naming convention: first number (1-99) changes with every change in temperature set point. Second number distinguishes each change in flow rate. =0 if only one flow rate tested at this temperature on the day. Respect the naming convention to be able to call data for plotting
Test Number | Date | Start Time | End Time | Temperature | Flow Rate | DL Range | Comment
| dd/mm | hh:mm | hh:mm | C | g/s | kW
10 | 2015-03-31 | 17:53 | 18:19 | -10 | up | 0 | Pure liquid, stepped flow
20 | 2015-04-01 | 9:20 | 10:15 | -10 | 70 | 0-7
30 | 2015-04-02 | 15:55 | 17:10 | -10 | 105 |  | poor accumulator management. Strange transient repsonse
40 | 2015-04-07 | 16:10 | 16:57 | -10 | 105 | 0-12
50 | 2015-04-08 | 10:47 | 12:08 | -10 | 70 | 0-8
61 | 2015-04-09 | 15:20 | 16:04 | -10 | 70 | 0-8.5
62 | 2015-04-09 | 16:18 | 17:00 | -10 | 30 | 2-7.0
63 | 2015-04-09 | 17:04 | 17:39 | -10 | 90 | 4-11.0
71 | 2015-04-10 | 15:08 | 15:44 | 5 | 70 | 2-7.0
72 | 2015-04-10 | 15:52 | 16:29 | 5 | 37.1 | 3-7.5 |  | was 30
81 | 2015-04-13 | 15:46 | 16:12 | -25 | 30 | 2-3.5 | Lost chiller several times. Abandoned test
82 | 2015-04-13 | 16:17 | 17:00 | -25 | 70 | 4-9.0
83 | 2015-04-14 | 17:13 | 17:37 | -25 | 105 | 4-8,0
91 | 2015-04-14 |  |  | 0 | 70
92 | 2015-04-14 |  |  | 0 | 105
93 | 2015-04-14 |  |  | 0 | 30
101 | 2015-04-15 |  |  | -10 | 50
102 | 2015-04-15 |  |  | -10 | 20
111 | 2015-04-17 |  |  | -10 | 54
112 | 2015-04-17 |  |  | -10 | 40
121 | 2015-04-20 |  |  | -10 | 60
122 | 2015-04-20 |  |  | -10 | 80
123 | 2015-04-20 |  |  | -10 | 40
131 | 2015-04-21 |  |  | -10 | 105
132 | 2015-04-21 |  |  | -10 | 50
133 | 2015-04-21 |  |  | -10 | 20
141 | 2015-04-22 |  |  | -10 | 90
142 | 2015-04-22 |  |  | -10 | 20
143 | 2015-04-22 |  |  | -10 | 40
144 | 2015-04-22 |  |  | -10 | 30
151 | 2015-04-23 |  |  | -20 | 70
152 | 2015-04-23 |  |  | -20 | 30
161 | 2015-04-23 |  |  | -5 | 70
162 | 2015-04-23 |  |  | -5 | 30
171 | 2015-04-23 |  |  | 12 | 70 | was 10
181 | 2015-04-24 |  |  | -10 | 90
182 | 2015-04-24 |  |  | -10 | 70
183 | 2015-04-24 |  |  | -10 | 60
184 | 2015-04-24 |  |  | -10 | 40

## Sheet2
2 Phase Test 1 Analysis 01/04 | [7 empty] | 2 Phase Test 1 Analysis 02/04
Time | 9:27 | 9:36 | 9:42 | 9:49 | 9:54 | 9:57 | 10:00 | Time | 15:56 | 16:01 | 16:07 | 16:11 | 16:15 | 16:21 | 16:25 | 16:29 | 16:31 | 16:34 | 16:41 | 16:46 | 16:52 | 16:58 | 16:58 | 17:11
hh:mm:ss | [7 empty] | hh:mm:ss
Heat Load | 0 | 2000 | 3000 | 4000 | 5000 | 6000 | 7000 | Heat Load | 0 | 2000 | 4000 | 6000 | 7000 | 8000 | 8500 | 9000 | 9500 | 10000 | 10500 | 11000 | 11500 | 12000 | 12500 | 13000
W | [7 empty] | W
PLANT | [7 empty] | PLANT
Reference Flow | 71.1 | 71.3 | 71 | 70.7 | 70.7 | 70.5 | 70.4 | Reference Flow | 104.8 | 104.1 | 103.3 | 103.3 | 103 | 101.8 | 102.3 |  |  | 101.9 |  |  | 101.5 |  |  | 100.5
FT3020 | [7 empty] | FT3020
g/s | [7 empty] | g/s
Plant Density | 1103 | 1113 | 1110 | 1110 | 1107 | 1107 | 1103 | Plant Density | 1110 | 1110 | 1107 | 1107 | 1110 | 1107 | 1105 |  |  | 1098 |  |  | 1095 |  |  | 1085
DT3020 | [7 empty] | DT3020
kg/m^3 | [7 empty] | kg/m^3
Plant Pressure | 27.8 | 30 | 31.2 | 31.4 | 31.6 | 32 | 32.4 | Plant Pressure
PT3012 | [7 empty] | PT3012
bar | [7 empty] | bar
Plant Temperature | -30.4 | -31.9 | -31.9 | -31.6 | -30.9 | -30.9 | -30.4 | Plant Temperature
TT3020 | [7 empty] | TT3020
C | [7 empty] | C
Predicted Plant Density | 1082 | 1090 | 1090 | 1089 | 1086 | 1086 | 1084 | Predicted Plant Density
kg/m^3 | [7 empty] | kg/m^3
Manifold Loop 7424 | [7 empty] | Manifold Loop 7424
Measured Flow | 72.6 | 70.5 | 63 | 53.8 | 39.1 | 24.4 | 16 | Measured Flow | 106.9 | 107 | 54.9 | 1.4 | 1.4 | 15 | 16.1 |  |  | 40.4 |  |  | 55.3 |  |  | 65.4
FT7524 | [7 empty] | FT7524
g/s | [7 empty] | g/s
Measured Density | 1026.6 | 458.1 | 270.5 | 202 | 170.3 | 154.5 | 148.6 | Measured Density | 1070 | 1009 | 270.5 | 235.4 | 235.4 | 241.2 | 235.4 |  |  | 235.4 |  |  | 216.6 |  |  | 199
DT7424 | [7 empty] | DT7424
kg/m^3 | [7 empty] | kg/m^3
Measured Temperature | -17.7 | -9.8 | -8.4 | -8.4 | -7.8 | -7.8 | -7.1 | Measured Temperature | -26.4 | -15.1 | -12 | -11.3 | -9.3 | -5.9 | -5.3 |  |  | -5.3 |  |  | -5.9 |  |  | -5.2
TT7424 | [7 empty] | TT7424
C | [7 empty] | C
Local Pressure at Instrument | 24.7 | 26.4 | 27.3 | 27.1 | 27.4 | 27.6 | 27.6 | Local Pressure at Instrument
PT7450 | [7 empty] | PT7450
bar | [7 empty] | bar
Manifold Return Temperature | -14.6 | -10.7 | -9.4 | -9.5 | -9.6 | -9.3 | -9.2 | Manifold Return Temperature | -23.4 | -15.1 | -14.4 | -14.6 | -12.9 | -9.3 | -8.2 |  |  | -9 |  |  | -9.9 |  |  | -9.9
TT7056 | [7 empty] | TT7056
C | [7 empty] | C
Manifold Supply Pressure | 27.4 | 29.3 | 30.6 | 30.7 | 31.1 | 31.4 | 31.8 | Manifold Supply Pressure
PT7024 | [7 empty] | PT7024
bar | [7 empty] | bar
Manifold Supply Temperature | -18.5 | -14.5 | -10.2 | -9.7 | -9.7 | -9.4 | -9.2 | Manifold Supply Temperature | -26.7 | -22.2 | -15.2 | -15 | -13.7 | -10 | -8.9 | [8 empty] | -10.2
TT7024 | [7 empty] | TT7024
C | [7 empty] | C
Predicted Density | 1028 | 523 | 341.5 | 268 | 229.5 | 198.4 | 173.7 | Predicted Density
Using h3 and heat load | [7 empty] | Using h3 and heat load
kg/m^3 | [7 empty] | kg/m^3
Flow Error | -1.5 | 0.8 | 8 | 16.9 | 31.6 | 46.1 | 54.4 | Flow Error | -2.1 | -2.9 | 48.4 | 101.9 | 101.6 | 86.8 | 86.2 | 0 | 0 | 61.5 | 0 | 0 | 46.2 | 0 | 0 | 35.1
g/s | [7 empty] | g/s
Relative Flow Error | -2.1097 | 1.122 | 11.2676 | 23.9038 | 44.6959 | 65.3901 | 77.2727 | Relative Flow Error | -2.0038 | -2.7858 | 46.8538 | 98.6447 | 98.6408 | 85.2652 | 84.262 | #DIV/0! | #DIV/0! | 60.3533 | #DIV/0! | #DIV/0! | 45.5172 | #DIV/0! | #DIV/0! | 34.9254
Loop 4 Pressure Drop | 2.7 | 2.9 | 3.3 | 3.6 | 3.7 | 3.8 | 4.2 | Loop 4 Pressure Drop | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
bar | [7 empty] | bar
2 Phase Density Deviation from prediction | -1.4 | -64.9 | -71 | -66 | -59.2 | -43.9 | -25.1 | 2 Phase Density Deviation from prediction | 1070 | 1009 | 270.5 | 235.4 | 235.4 | 241.2 | 235.4 | 0 | 0 | 235.4 | 0 | 0 | 216.6 | 0 | 0 | 199
Relative Deviation | -0.13619 | -12.4092 | -20.7906 | -24.6269 | -25.7952 | -22.127 | -14.4502 | Relative Deviation | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Vapour Quality in Loop 7424, predicted from density measurement and local pressure | 0 | 0.0891 | 0.213 | 0.31 | 0.388 | 0.4396 | 0.4603 | Vapour Quality in Loop 7424, predicted from density measurement and local pressure | 0 | 0.0891 | 0.213 | 0.31 | 0.388 | 0.4396 | 0.4603 | 0.0891 | 0.213 | 0.31 | 0.388 | 0.4396 | 0.4603 | 0.388 | 0.4396 | 0.4603
Flow Regime | [7 empty] | Flow Regime

## Sheet3
Test Number | Date | Start Time | End Time | Temperature | Flow Rate | Test Number | Date | Start Time | End Time | Temperature | Flow Rate
| dd/mm | hh:mm | hh:mm | C | g/s |  | dd/mm | hh:mm | hh:mm | C | g/s
10 | 2015-03-31 | 17:53 | 18:19 | -10 | up | 121 | 2015-04-20 |  |  | -10 | 60
20 | 2015-04-01 | 9:20 | 10:15 | -10 | 70 | 122 | 2015-04-20 |  |  | -10 | 80
30 | 2015-04-02 | 15:55 | 17:10 | -10 | 105 | 123 | 2015-04-20 |  |  | -10 | 40
40 | 2015-04-07 | 16:10 | 16:57 | -10 | 105 | 131 | 2015-04-21 |  |  | -10 | 105
50 | 2015-04-08 | 10:47 | 12:08 | -10 | 70 | 132 | 2015-04-21 |  |  | -10 | 50
61 | 2015-04-09 | 15:20 | 16:04 | -10 | 70 | 133 | 2015-04-21 |  |  | -10 | 20
62 | 2015-04-09 | 16:18 | 17:00 | -10 | 30 | 141 | 2015-04-22 |  |  | -10 | 90
63 | 2015-04-09 | 17:04 | 17:39 | -10 | 90 | 142 | 2015-04-22 |  |  | -10 | 20
71 | 2015-04-10 | 15:08 | 15:44 | 5 | 70 | 143 | 2015-04-22 |  |  | -10 | 40
72 | 2015-04-10 | 15:52 | 16:29 | 5 | 37.1 | 144 | 2015-04-22 |  |  | -10 | 30
81 | 2015-04-13 | 15:46 | 16:12 | -25 | 30 | 151 | 2015-04-23 |  |  | -20 | 70
82 | 2015-04-13 | 16:17 | 17:00 | -25 | 70 | 152 | 2015-04-23 |  |  | -20 | 30
83 | 2015-04-14 | 17:13 | 17:37 | -25 | 105 | 161 | 2015-04-23 |  |  | -5 | 70
91 | 2015-04-14 |  |  | 0 | 70 | 162 | 2015-04-23 |  |  | -5 | 30
92 | 2015-04-14 |  |  | 0 | 105 | 171 | 2015-04-23 |  |  | 12 | 70
93 | 2015-04-14 |  |  | 0 | 30 | 181 | 2015-04-24 |  |  | -10 | 90
101 | 2015-04-15 |  |  | -10 | 50 | 182 | 2015-04-24 |  |  | -10 | 70
102 | 2015-04-15 |  |  | -10 | 20 | 183 | 2015-04-24 |  |  | -10 | 60
111 | 2015-04-17 |  |  | -10 | 54 | 184 | 2015-04-24 |  |  | -10 | 40
112 | 2015-04-17 |  |  | -10 | 40